## Sheet1
Name | Centrifuges, Tabletop
Code | E-LA-255
Follower/Leader | Follower
Technology Level | Secondary/Tertiary
Purpose | Centrifuges use centrifugal force to separate suspended particles from a liquid or to separate liquids of different densities.
Specialty | Clinical laboratories
General Specifications | Mobility | Yes
| FDA Clearance | Yes
| CE Mark | Yes
| Portability | Yes
| Weight, kg (lb) | 128 (283)
| Dimentions (H x W x D), cm (in) | 34.3 x 76.2 x 62.2 (13.5 x 30 x 24.5)
Technical Specifications | Max rpm | 10200
| Max RCF, g | 11400
| Speed adjustment | Microprocessor controlled, manual in 10 RPM increments
| Max accel time, sec | 90
| Rotor Identification | Yes
| Maximum Capacity, mL | 3,000 (4 x 750)
| Temperature Range, °C | Ambient control (4 to 35)
| Brush/Brushless | Brushless Induction
| Lid interlock | Yes
ALERT INDICATORS
| Imbalance | Yes
| End of run | Yes
| Overspeed | Yes
| Other | Lid open, error codes, rotor recognition
| Display | NA
TIMER
| Range, min | 0-60, pulse
| Settings | 1 min increments
| Accuracy | <0.05 sec